# GENERAL

## SECTION INCLUDES

### Material and installation criteria for Ceiling Lifts installed in Education Facilities.

## RELATED SECTIONS

### Section 01 33 00 – Submittal Procedures.

### Section 01 61 00 – Common Product Requirements.

### Section 01 78 00 – Closeout Submittals.

### Section 05 50 00 – Metal Fabrications.

## REFERENCES

### Canadian Standards Association (CSA)

#### CSA C22.2 No. 60601-1, Medical Electrical Equipment – Part 1: General Requirement for Basic Safety and Essential Performance.

### Underwriters Laboratories (UL)

#### UL 60601-1, Medical Electrical Equipment: General Requirements for Safety.

## SUBMITTALS

### Indicate size and description of components, base material, structural support and electrical connections.

### Shop drawings for structural support track system to be stamped by structural engineer licensed to practice in the Province of Newfoundland and Labrador.

## CLOSEOUT SUBMITTALS

### Provide operation and maintenance data for incorporation into manual specified in Section 01 78 00 – Closeout Submittals.

### Deliver ceiling lift equipment only after interior finishes are complete.

# PRODUCTS

## COMPONENTS

### Ceiling lift system for transfer of physically challenged people to and from wheelchairs, toilets and showers.

### Features:

#### Lifting capacity: 272 kg.

#### LED indicator for maintenance required.

#### Electronic microprocessor soft-start and stop motor control.

#### Power on indicator.

#### Double centrifugal emergency brake system in case of mechanical failure.

#### Manual emergency lowering device.

#### Electrical up and down emergency buttons.

#### Emergency stop pull cord accessible from the ground.

#### Overload circuit protection current limiter.

#### Low battery disconnect system.

#### Audible and visual (LED) low battery indicator.

#### Changing indicators: changing and changing complete.

#### Strap length: up to 2.3 m tested for 1360 kg.

#### Lifting speed: 2.7 cm/s at 272 kg.

#### Batteries: 2 x 5Ah averaging 150 cycles at 75 kg loading.

#### Adjustable horizontal displacement speeds: 10, 15, 20 and 25 cm/s.

#### Automatic return to change function.

#### Fire resistant ABS casing.

#### Certification: CSA 22.2 No. 60601-1, UL No. 60601-1.

#### CE marked.

#### Electromagnetic interference: to EMI standards.

#### ISOS 10535.

#### Charger unit:

##### Power indicator on charging module.

##### Clip on charger anywhere on track.

##### 120/240 VAC/60 Hz/27 Va.

##### Class 2 insulated.

#### Hand control:

##### Fire retardant ABS.

##### Santropen tactile buttons.

##### IP44.

#### Track configuration:

##### Manufacturer’s standard track system. Constructed of extruded aluminum. Track system comprising of straight track, curves and connection features.

###### Straight track sections in 90 mm, 140 mm and 180 mm heights of longest practical lengths. Manufacturer to determine track height based on spacing of structural support.

##### Curve sections available 45 and 90 degrees.

##### Installation to include all required threaded rods, ceiling plates, shims, track brackets with locks, Hilti struts, steel plates, lateral bracing, nuts, bolts and washers.

##### Track configuration to include all required gates, exchanger and turntable components.

# EXECUTION

## EXAMINATION

### Take critical site dimensions to ensue tolerances and clearances to other constructions have been maintained, and necessary adjustments are made to adapt work of this section during its installation.

### Ensure anchors, services and similar provisions, installed by others, are adequate to meet specified requirements, and make adjustments before installation.

### Do not proceed with work of this section until conditions and work on which it depends are satisfactory with limitations of adjustment.

## INSTALLATION

### Install in accordance with manufacturer’s instructions and regulations in force.

## PROTECTION

### Provide protective coverings for finished surfaces.

## TOUCH-UP

### Upon completion, touch-up and restore to new condition, damaged or defaced factory finished surfaces.

### Remove protective coverings and clean exposed surfaces after completion.

## FIELD QUALITY CONTROL

### Engage only specialized tradesmen employed and under direct supervision of Ceiling Lift manufacturer.

### Ensure Ceiling Lift operation is free of vibration and performs smoothly and quietly.

## COMMISSIONING

### Train user staff in operation, cleaning and maintenance of ceiling lift equipment provided under this section.

END OF SECTION